## Slide 1
Development of cleaning agents for biofilm removal
Philipp Stiefel1, Stefan Mauerhofer2, Stefanie Altenried1, Qun Ren1, Urs Rosenberg2
1) Laboratory for Biomaterials, Swiss Federal Laboratories for Materials Science and Technology (Empa), Lerchenfeldstrasse 5, CH-9014 St.Gallen, Switzerland. 2) Borer Chemie AG, Gewerbestrasse 13, CH-4528 Zuchwil, Switzerland
Sponsored by:

## Slide 2: Biofilms
7-10 oct. 2015
Structured communities of microorganisms on surfaces or at interfaces.
Self-made matrix made from extracellular polymeric substances (EPS) consisting of polysaccharides, proteins, lipids, nucleic acids and oder substances.
Biofilm-associated microorganisms up to 1000 fold less sensitive to antimicrobial substances compared to their planctonic forms.
1

## Slide 3
Endoscopes and infections
Kovaleva et al. 2013
7-10 oct. 2015
2

## Slide 4: Biofilm in endoscope channels
7-10 oct. 2015
REM-pictures of two different air-/water channels with biofilm:
confluent soil-biofilm-layer (low resolution).
multi-layer biofilm composed of healthy appearing cells surrounded by and overlaid with amorphous appearing EPS (high resolution).
Pajkos et al. 2004
3

## Slide 5: Cleaning models
7-10 oct. 2015
Screening-methods for testing large numbers of formulations and parameters plus different methods for analysis:  microtiter plate format
More laborious methods closer to the real application of products:  flow cell  endoscope channels (EN ISO TS 15883-5)  others?
4

## Slide 6: Growing biofilms in microtiter plates
7-10 oct. 2015
96-well microtiter plates from polystyrene with flat bottom.
Growing biofilms: dilute pre-culture to 0.2 OD in 30% TSB + 0.25% glucose. Incubate 200 L / well at 33 °C while shaking at 40 rpm during 24 hours.
Appearance of biofilms (crystal violet staining):
P. aeruginosa
S. aureus
5

## Slide 7: Cleaning experiments in microplates
7-10 oct. 2015
Sucking off medium
Washing 1x with 300 l 0.9% NaCl
Adding of 300 l cleaner sol. (1% in water of standardized hardness (WSH))
Incubating for 40 min at 25°C without shaking
Removing cleaner solution
Wash 3x with 300 l 0.9% NaCl
Applying analysis method
6

## Slide 8: Potential analysis methods
7-10 oct. 2015
Total biomass (bacteria and EPS)  Crystal Violet (CV)  Safranin Red  Congo Red
Total bacteria (live and dead)  SYTO 9  Acridine Orange
Living bacteria  SYTO 9 / Propidium Iodide  BacTiter-GloTM assay (ATP)  Turbidity Threshold (growth kinetics)
Proteins  CBQCA  NanoOrange  SYPRO Ruby  FITC
Polysaccharides  Calcofluor White  ConA-FITC  Dubois
7

## Slide 9: Advantages / disadvantages
7-10 oct. 2015
Application in microtiter plates
8

## Slide 10: Assays – what they can tell us
7-10 oct. 2015
Crystal Violet stains total biomass (bacteria snd EPS)  biofilm removal without differentiation between EPS and bacteria.
BacTiter-GloTM differential staining of living bacteria  bacteria removal and/or –killing without differentiation of the two.
Turbidity Threshold bacterial growth kinetics (OD595)  bacteria removal and/or –killing without differentiation of the two.
9

## Slide 11: P. aeruginosa / S. aureus biofilm removal
7-10 oct. 2015
Crystal Violet (CV) staining
S. aureus
P. aeruginosa
Optical density
Optical density
10

## Slide 12: Incubation time and biofilm removal
7-10 oct. 2015
P. aeruginosa (CV-staining)
Residue [% of NaCl]
11

## Slide 13: Incubation temperature and biofilm removal
7-10 oct. 2015
Residue [% of NaCl]
P. aeruginosa (CV-staining)
12

## Slide 14: Growth of remaining bacterial Turbidity
7-10 oct. 2015
RF 1.48 log
RF 1.84 log
RF 5.98 log
13
Serial dilution for calibration
Turbidity Threshold

## Slide 15: Products under scrutiny
7-10 oct. 2015
14

## Slide 16: TOSI-cleaning
7-10 oct. 2015
R1-m
DR-m
R3-m
NP-m
R6-o
PosK
WPO
0 min
30 min
15 min
45 min
60 min
rinsed dried
1% in tap water of ca. 16°d at room temperature (ca. 22°C)
15

## Slide 17: P. aeruginosa vs. S. aureus
7-10 oct. 2015
Microtiter plate, biofilm reduction 40 min static conditions, CV-staining
287
134
16

## Slide 18: Living cells reduction (P. aeruginosa)
7-10 oct. 2015
Microtiter plate, 40 min static conditions, BacTiter-GloTM assay
17

## Slide 19: Biofilm & living cells reduction
7-10 oct. 2015
18
Crystal Violet & BacTiter-GloTM assay

## Slide 20: Flow cell with OCT-visualization
7-10 oct. 2015
Miniaturized flow cell 120x2x1 mm (LxWxH)
Heterothrophic waste water biofilm (source: supernatant of activated sludge) cultivated during 24h in fed-batch mode.
Rinsed with 0.9% NaCl
Cleaning experiments in flow through mode (16 ml/min)
Experiments with NP-m and PosK.
Optical coherence tomography (OCT) is an established medical imaging technique that uses light to capture micrometer-resolution, 3d images from within optical scattering media (e.g., biological tissue). OTC is based on low-coherence interferometry, typically employing near-infrared light. The use of relatively long wavelength light allows it to penetrate into the scattering medium (Wikipedia).
19
Time resolved video of 2d-sections (xz-plane) along flow path (x-direction)

## Slide 22: Endoscope dummy test
7-10 oct. 2015
Filling a 2 m endoscope channel  4 mm with 10 ml P. aeruginosa over night culture.
Growing biofilm for 72 h at 30°C by recirculation (40 mL/min) with defined addition of growth medium.
Rinsing channel with 0.9% NaCl for 1 min at 20 ml/min.
Cutting channel in 30 cm sections for cleaning experiments (static or with flow through at 200 ml/min (cleaning) and 4 ml/min (disinfection)).
Analysing for OD595nm and colony forming units (CFU) and carrying out quantification of sugar (Dubois) and protein (Lowry).
EN ISO/TS 15883-5:2005 Appendix F
20

## Slide 23: E. channel cleaning - experimental setup
7-10 oct. 2015
treatment
Rinsing
0.9% NaCl
20ml/min
1min
Cleaning
1% cleaner
Contr.: WSH
200ml/min
5 / 15min
Vortex
5min
Cutting
longitudinal in 2 halves
5mm sections
Disinfection
deconex® HLD PA / PA20
4ml/min
15min
Rinsing
0.9% NaCl
20ml/min
1min
Rinsing
0.9% NaCl
20ml/min
1min
Without disinfection
21

## Slide 24: Analysis methods
7-10 oct. 2015
120µl
120µl
Sample
150µl
30µl
30µl
120µl
30µl
…
Serial dilution
(bacteria)
Lowry method (protein)
Dubois method
(sugar)
22
OD595 (bacteria)

## Slide 25: Bacteria reduction
7-10 oct. 2015
1.6
2.5
2.1
RF log10
0.8
Static (without pumping) vs. flow through cleaning (pumping at 200 ml/min)
23

## Slide 26: Reduction of protein and sugar
7-10 oct. 2015
Flow through cleaning: 5, 15 min, 200 ml/min
24
NP-m fulfills requirements of EN ISO/TS 15883-5:2005 Appendix F

## Slide 27: Summary reduction of all parameters
7-10 oct. 2015
25
Flow through cleaning:15 min, 200 ml/min

## Slide 28: Comparison dummy / microplate
7-10 oct. 2015
 Dummy, flow through
 Microplate
26

## Slide 29: Residual biofilm under microscope
7-10 oct. 2015
WSH
NP-m
R2-m
R3-m
R1-m
R6-o
250 m
250 m
250 m
250 m
250 m
250 m
CYTO 9-staining (nucleic acids)
27

## Slide 30: Residual biofilm under microscope
7-10 oct. 2015
WSH
NP-m
R2-m
R3-m
R1-m
R6-o
25 m
25 m
25 m
25 m
25 m
25 m
28
Higher resolution

## Slide 31: Without / with disinfection
7-10 oct. 2015
Cleaning 15 min, 200 ml/min  disinfection 15 min, 4 ml/min
39

## Slide 32: Biofilm before / after disinfection
7-10 oct. 2015
250 m
250 m
25 m
25 m
25 m
25 m
WSH cleaning
WSH cleaning
After WSH & Disinfection
After WSH & Disinfection
After NP-m & Disinfection
NP-m cleaning
Disinfectant: deconex® HLD PA (formulated pH-neutral peracetic acid)
30

## Slide 33: Summary
7-10 oct. 2015
Different analysis methods for different phases of product development required.
Results of microplate screening methods essentially congruent with more laborious dummy tests.
S. aureus biofilm less resistant than P. aeruginosa biofilm, particularly when using enzymatic cleaners.
Disinfectant-cleaners and cleaners with biocidal ingredients generally didn’t clean effectively, neither biofilm nor blood. They showed particularly low performance in case of S. aureus biofilm.
Only excellent cleaning guarantees the success of the following disinfection stage. This is especially true in case of biofilm.
There is often a big gap between manufacturer claims and real performance.  Need for standardized test methods and definition of minimal performance requirements for medical device cleaners.
Within this project, a new enzymatic product with excellent cleaning performance for biofilm and blood was developed, which will be used for manual (pre-)cleaning of flexible endoscopes as well as surgical instruments.
31

## Slide 34: Thank you for your attention!
7-10 oct. 2015
Dr. Qun Ren, Empa
Stefan Mauerhofer, Borer
Dr. Philipp Stiefel, Empa
Dr. Urs Rosenberg, Borer
Stefanie Altenried, Empa
Jana Schneider, Empa
The project team
32